## Sheet1
mp2rage Slab
| TA |  | 9:07
| PAT |  | 2
| Voxel size |  | 0.6x0.6x0.6mm
| Rel. SNR |  | 1.0
Properties
| Prio Recon |  | Off
| Load to viewer |  | On
| Inline movie |  | Off
| Auto store images |  | On
| Load to stamp segments |  | Off
| Load images to graphic segments |  | Off
| Auto open inline display |  | Off
| Start measurement without further preparation |  | On
| Wait for user to start |  | Off
| Start measurements |  | single
Routine
Slab group 1
|  | Slabs | 1
|  | Dist. factor | 50%
|  | Position | R2.4 A29.1 H23.0
|  | Orientation | T>C-23.0
|  | Phase enc. dir. | R>>L
|  | Rotation | 90.00 deg
| Phase oversampling |  | 0%
| Slice oversampling |  | 100.0 %
| Slices per slab |  | 128
| FoV read |  | 192 mm
| FoV phase |  | 81.3%
| Slice thickness |  | 0.60 mm
| TR |  | 5000 ms
| TE |  | 3.71 ms
| Averages |  | 1
| Concatenations |  | 1
| Filter |  | None
| Coil elements |  | A24
Contrast
| Magn. preparation |  | Non-sel. IR
| TI 1 |  | 900 ms
| TI 2 |  | 2750 ms
| Flip angle 1 |  | 5 deg
| Flip angle 2 |  | 3 deg
| Fat suppr. |  | None
| Water suppr. |  | None
| 2nd Inversion-Contrast |  | On
| Averaging mode |  | Long term
| Reconstruction |  | Magnitiude
| Measurements |  | 1
| Multiple series |  | Each measurement
Resolution
| Base resolution |  | 320
| Phase resolution |  | 100%
| Slice resolution |  | 100%
| Phase partial Fourier |  | 6/8
| Slice partial Fourier |  | 6/8
| Interpolation |  | Off
| PAT mode |  | GRAPPA
| Accel. factor PE |  | 2
| Ref. lines PE |  | 24
| Accel. factor 3D |  | 1
| Reference scan mode |  | Integrated
| Image Filter |  | Off
| Distortion Corr. |  | Off
| Prescan Normalize |  | Off
| Normalize |  | Off
| B1 filter |  | Off
| Raw filter |  | Off
| Elliptical filter |  | Off
Geometry
| Multi-slice mode |  | Single shot
| Series |  | Interleaved
| Table position |  | H
| Table position |  | 0 mm
| Inline Composing |  | Off
System
| V24 |  | Off
| A24 |  | On
| Positioning mode |  | FIX
| MSMA |  | S-C-T
| Sagittal |  | R >>L
| Coronal |  | A>>P
| Transversal |  | F>>H
| Save uncombined |  | Off
| Coil Combine Mode |  | Adaptive Combine
| AutoAlign |  | ---
| Auto Coil Select |  | Off
| Shim mode |  | Standard
| Adjust with body coil |  | Off
| Confirm freq. adjustment |  | Off
| Assume Silicone |  | Off
| ? Ref. amplitude 1H |  | 0.000 V
| Adjustment Tolerance |  | Auto
Adjust volume
|  | Position | R2.4 A29.1 H23.0
|  | Orientation | T>C-23.0
|  | Rotation | 90.00 deg
|  | A>>P | 192 mm
|  | R>>L | 156 mm
|  | F>>H | 77 mm
Physio
| 1st Signal/Mode |  | None
| Dark blood |  | Off
| Resp. control |  | Off
Inline
| Subtract |  | Off
| Std-Dev-Sag |  | Off
| Std-Dev-Cor |  | Off
| Std-Dev-Tra |  | Off
| Std-Dev-Time |  | Off
| MIP-Sag |  | Off
| MIP-Cor |  | Off
| MIP-Tra |  | Off
| MIP-Time |  | Off
| Save original images |  | On
Sequence
| Introduction |  | On
| Dimension |  | 3D
| Elliptical scanning |  | Off
| Asymmetric echo |  | Allowed
| Contrasts |  | 1
| Bandwidth |  | 240 Hz/Px
| Flow comp. |  | Slice
| Echo spacing |  | 7.5 ms
| RF pulse type |  | Fast
| Gradinet mode |  | Whisper
| Excitation |  | Slab-sel.
| RF spoiling |  | On
| FFT Scale Factor |  | 100%
| Line/Partition Swap |  | Off
| Homodyne Phase Filter |  | Off
| Flat Image |  | On
| T1 Map |  | On
| Division Image |  | Off
| ExtInvPulseOn |  | On
| OffResFreqInv |  | 0
| Invflipangle |  | 1800