/*!

\page qatools QA tools

This package provides several tools to perform QA (quality assurance)

calculations.  They produce images, graphs, and/or XML data as output.

<tt>fmriqa_phantomqa.pl</tt>

and <tt>fmriqa_generate.pl</tt>

are the main tools and produce an HTML report with various QA measures.

<tt>fmriqa_phantomqa.pl</tt>

was designed for fMRI images of the BIRN stability phantom, and

<tt>fmriqa_generate.pl</tt> has been used for human fMRI data.

\section phantomqa BIRN Agar Phantom QA

The tool <tt>fmriqa_phantomqa.pl</tt> implements several recommendations of the

fBIRN Calibration working group in analyzing fMRI data from a spherical

agar-filled phantom.  A subset of these recommendations are

implemented as described in the following publication:

Friedman, L., Glover, G., "Report on a Multicenter fMRI Quality Assurance Protocol", J Magn Reson Imaging.  June 2006.  23(6):827-39.

Notes are included below for any metric whose implementation differs from

that described in the report.

To run the tool, provide a minimum of two arguments -- first, the image data as

wrapped by a BXH or XCEDE file, and second, the name of the directory where

the outputs should go.  This directory should not exist, unless the --overwrite

option is specified.  So, for example, if you have a set of DICOM files, you

could do something like:

\code

dicom2bxh *.dcm WRAPPED.bxh

fmriqa_phantomqa.pl WRAPPED.bxh OUTPUTQADIR

\endcode

Several options to the QA tool are available, and documentation of these are

available by running <tt>fmriqa_phantomqa.pl --help</tt>.

To avoid cutting and pasting wholesale from the text of the paper, we refer to the Friedman, Glover (2006) paper for documentation and justifications of several metrics.

The differences between (and additions to) the metrics described in the paper are described below:

<ul>

<li><b>Initially discarded volumes</b>: if the number of timepoints in the data is even, then the first 2 timepoints are discarded, as recommended in the paper.  However, since several metrics depend on an even number of volumes, if the number of timepoints is odd, then 3 timepoints are discarded.

<li><b>ROI</b>: the ROI used in these computation (when specified) is an NxN square in the center of the middle slice of the volume.  This tool uses N=15 voxels for a 64x64 voxel image by default, but can be adjusted with the --roisize option.  The current recommendation from the paper for a 64x64 voxel image is an ROI of 21x21 voxels, but the default remains at 15x15 to maintain compatibility with earlier revisions of the protocol.</li>

<li><b>Drift</b>: in the paper, the signal <i>drift</i> is calculated on the second-order polynomial fit of a signal composed of the mean intensities for each volume across the ROI.  This value is listed as <tt>driftfit</tt> in the output of this QA tool.  The value <tt>drift</tt> as calculated by this tool, however, is calculated on the unfitted (i.e. raw) signal.</li>

<li><b>Fourier analysis</b>: the y-axis in the frequency spectrum images generated by this tool is scaled by the mean of the raw mean ROI signal (see above), and matches what the paper suggests will be implemented in a "future version" of the protocol.  To limit any confusion, the y-axis of the plot is annotated as "mean scaled".</li>

<li><b>Additions</b>: several newer analyses beyond those described in the paper have been added to the QA report (and are documented below).</li>

</ul>

The GSL function <tt>gsl_multifit_linear()</tt> is used to perform the polynomial fit.

The output directory will contain several images and data files, including an XML file 'summaryQA.xml' which encodes many of the QA metrics into XCEDE2 format.

The images and several acquisition and summary values can be displayed by opening the file <tt>index.html</tt> in any web browser.

At the top of index.html are a series of images based on recommendations from the Friedman, Glover paper:

<ul>

<li><i>Raw signal</i>: this image plots the mean intensity of the raw signal across the ROI for each timepoint, as well as the second-order polynomial fit of the same signal.  Displayed in the graph are the summary values <i>percent fluctuation</i>, <i>drift</i> and <i>driftfit</i> (see note above).  These values are also available in <tt>summaryQA.xml</tt> as "percentFluc", "drift", and "driftfit".</li>

<li><i>Magnitude spectrum</i>: this image plots a Fourier transform of the signal, with the y-axis scaled to be a percentage of the mean of the raw signal intensity across the ROI.  Some other summary values are also displayed here: "mean", "SNR", and "SFNR" (also in <tt>summaryQA.xml</tt> with the same names).</li>

<li><i>Weiskoff analysis</i>: this plots the observed and theoretical coefficients of variation for voxels within various ROI sizes up to the maximum ROI size.

</ul>

Several more images follow:

<ul>

<li><i>FWHM (full-width half-maximum)</i>: if AFNI tools were installed when running the QA, three images plotting the FWHM values over time (in each of X, Y, and Z dimensions).  FWHM is calculated on the motion-corrected, 2nd-order polynomail detrended, masked data.  The order of operations follows this sequence (note that 2 or 3 volumes [by default] are removed from the input before executing):

<code>

3dvolreg -prefix REG INPUT

3dDetrend -polort 2 -prefix DETREND REG

3dTstat -mean -prefix MEAN REG

3dAutomask -q -prefix MASK MEAN

3dFWHMx -dset DETREND -mask MASK -out FWHMVALS

</code></li>

<li><i>Center of Mass</i>: the center of mass in the X, Y, or Z dimension is the weighted mean intensity over each volume, where the weights are the X, Y, or Z indices of each voxel.  If the phantom is always fully within the field of view, this can be a simple indication of whether the object being imaged somehow changed position during scanning.  If AFNI tools were installed when running the QA, the more sophisticated output of 3dvolreg is also plotted here (shifted along the y-axis to match the first value of the center-of-mass).</li>

<li><i>Ghost analysis</i>: for EPI sequences, this image plots the "ghostiness" of the data.  The ghost metric is calculated for each volume by calculating a dilated mask ("original mask") of the motion-corrected, detrended data, and shifting it by N/2 voxels in the appropriate axis to create a "ghost mask".  The mean intensities of those voxels in the original mask and not in the ghost mask, and of voxels in the ghost mask and not in the original mask are calculated.  The mean intensity of the top 10 percent of ghost-only voxels ("meanBrightGhost") is also calculated.  The ghost values are plotted as a percentage of the mean intensity of non-ghost voxels.  The dilated mask is generated using the following sequence of AFNI commands:

<code>

3dvolreg -prefix REG INPUT

3dDetrend -polort 2 -prefix DETREND REG

3dTstat -mean -prefix MEAN REG

3dAutomask -q -dilate 4 -prefix MASK MEAN

</code></li>

</ul>

Four more images from the original Friedman, Glover paper follow:

<ul>

<li><i>Odd-even difference</i>: this corresponds to the "Static Spatial Noise Image" in the paper.</li>

<li><i>Mean</i>: this corresponds to the "Signal Image" in the paper.</li>

<li><i>Standard Deviation</i>: this corresponds to the "Temporal Fluctuation Noise Image" in the paper.</li>

<li><i>SFNR</i>: this corresponds to the "Signal-toFluctuation-Noise Ratio (SFNR) Image" in the paper.</li>

</ul>

After this follows a table of the acquisition parameters, if available, extracted from the original input data files.

*/